OLIS(MCIS)
 PV1/Water/Vacuum Interlock or TWT Failure Response
(based on emails & conversations with Keerthi)
If beam disappears from MCIS, and one notes that RF Amplitude (dB) is at -20db, this likely indicates a water interlock or MCIS:IG1 trip (or PV1 trip). To restore beam, first clear interlock conditions (ensure waterflow OK, IG1 reset, Valves open), then return amplitude to its previous setting.

At this stage, it is also a good idea to check MCIS Optics and ensure both MCIS:BIAS0 & EL1 haven't also tripped. In the event that they did, their readbacks will be less than 100V. Resetting them is done by opening both the BIAS0 & EL pages, and clicking on the top right INTBIAS button. From there, the bias may be reset once the interlock condition is cleared.
Once certain that BIAS0/EL1 are back up to setpoint, check IOS:FC3 for beam.

If none can be found, the next step is to verify the status of the Travelling Wave Tube (TWT), into which the frequency generator connects. It is possible that it trips off due to a number of conditions, including high reflected RF (which can either occur for poor vacuum or rapid fluctuations in RF amplitude). The TWT control interface can be accessed by web browser at 142.90.120.72 The username and password are located in the ICR passwords list.

Note – there are no readbacks on the local unit. There is no longer a need to perform a local inspection to reset, everything can be done from the control room.

The fault can be cleared by clicking RESET, followed by TRANSMIT, on the right hand side menu, which should make set the unit back into operation. Note that TWT is set to trip off at 20W reflected power, to protect the unit from damage.

Next, revisit the RF Gen & Amp page. Try scanning RF Frequency around by a few GHz (click & drag). Keep an eye on FC3 – you should see beam appear. If that doesn't work, try increasing the RF amplitude (by small steps – less than 1dB – and keep an eye on the reflected RF) and repeat. As a reference, this procedure was successfully carried out with RF Amplitude at 0.92dB. Trying to rail the amplitude will only trip the unit off again.