DTL 4 manual start up:
1. Set setpoint to roughly 250
2. Set amplitude limit to roughly 1000
3. RF ON
4. Pulse mode
5. Set open loop amplitude drive to 350-400
6. Change to CW mode
7. Set tuner drive to -30 and hit RUN
8. Watch the reverse power bar. If the reverse power decreases wait until it is 20 or less then close tuning loop.
9. If the reverse power in 8. Increases change the tuner drive to +30 hit RUN and redo step 8.
10. After the tuner loop is closed, close Amplitude Loop
11. Then close Phase Loop
12. Increase setpoint and limit. Note: Increase setpoint slowly the cavity warms up slowly due to RF heating too.
13. When operated with high power, the reflected power will start oscillating after about an hour and settle down after about 3 hours (the cavity coupling can change in the first three hours of operation)