SCRF auto-on fail response – P/S rm

VZ, July 6, 2014

Always ensure RF is off before turning on amplifier.

- Insert coupler all the way.

- Open Loop Drive ~500.

- RF in CW.

- RF on.

- Look for amplitude readback ~100.  If not present, toggle pulse/CW.  If this fails, leave in pulse for ~10min. then try again.

- Check forward power (Pfwd) & adjust for ~150/70W SCB/SCC using open loop drive.

- Withdraw coupler until amplitude readback ~matches setpoint.

- Lock amplitude.

- Adjust phase rotator to minimize Pfwd.

- Adjust tuner setpoint so that tuner error fluctuates around '0.'

- Align tuner.

- Once frequency error dF is ~2Hz, 'auto-on.'

*note:  for SCB2:CAV4, Pfwd must be 120W locally (reads ~200W in EPICS).